## Slide 1: Meaningful G/T Measurements – Made at Night
International Telemetering Conference
October 2023

## Slide 2: Why Do We Care About G/T?
Short-term checkup
  Is it the same as last week, last month…
  This does not require an accurate absolute value
  Relative to previous measurements is all we need
Long-term planning
  G/T is an additive term in the link budget
  If it’s short of the predicted value, your link margin will be short
  This requires a “real” value
How do we get it?

## Slide 3: Solar-Noise G/T Method
Point antenna at the sun, get Ph in dBm
Point at cold sky, get Pc in dBm
YdB = Ph - Pc
Calculate Y factor: Y = 10^(YdB/10)
Estimate Fudge Factors
  C = Beam correction factor
  A = Atmospheric attenuation factor

## Slide 4: Solar Flux Uncertainties
Solar flux (S0) varies with
  Location (only measured in a few places)
  Frequency (only measured at a few frequencies)
    No station measures flux at telemetry frequencies
  Date and time (multiple cyclic variations)
    Carrington rotation every ~27 days
Interpolation algorithms needed
No “certain” value

## Slide 5: Y Factor Uncertainties
Y factor varies
  Inherently noisy measurement
  Cold sky temperature varies
    Sidelobes pick up interference
    Low elevations see more Earth, raising temperature
  Environmental noise raises noise floor
  Antenna tower height affects background noise
  Almost unmeasurable for small antennas
Standard procedure uses “coldest sky”
  Deliberately creates an optimistic value
  Not the value you should expect during operations

## Slide 6: Solar G/T Shortfalls
Result is not repeatable
  ±1 dB common from day to day
  Larger variations from place to place
Result is deliberately biased to “look good”
  Actual effective G/T may be 3 to 4 dB worse
Can’t be used to characterize multi-beam antennas
  Move to a planet with multiple suns?
Surely, we can do better…

## Slide 7: Replacing the Sun
Sun is a noise source
Use a signal source instead, with…
  Calibrated EIRP (variable power)
  High fidelity modulation (variable bit rate)
  Polarization control (linear or circular)
  Located at a calibrated distance
  Power adjusted to achieve a known Eb/N0
    Example: 11.4 dB in SOQPSK mode, where BER = 10-5
Start with a basic link budget equation:

## Slide 8: A Little Math Exercise…
  10 log () = EIRP, dBm
  10 log = Lp, Path loss, dB
  10 log = ( Polarization loss, dB
  10 log () = -198.6, dBm/(K-Hz)
  10 log = G/T, dB/K
  10 log () = Bit rate, dB-Hz

## Slide 9: Leading to…
We know everything on the right
  Eb/N0 can be determined from the known BER performance of the receiver
    Or DQM if it’s accurate
  Calibrated signal source gives us known EIRP
  Polarization control sets the polarization loss
  Measuring tape (or Google maps) gives us path loss
    Might need some math to account for height
  We know the data rate
We can get G/T with no reference to the sun

## Slide 10: What Do We Need?
Calibrated EIRP source
Known polarization
  Better yet, programmable polarization
Calibrated receiver
  Known BER vs. Eb/N0 performance
Calibrated distance
  Into the far field, at least
Line of sight path
Accurate pointing
  Cameras can help

## Slide 11: Field Test Setup
QSight™ on a 125-foot mast
PD450 with 6-foot dish

## Slide 12: QSightTM L/S/C-BAND BORESIGHT SYSTEM

## Slide 13: Typical result from verification testing
Frequency, Power Level, Polarization

## Slide 14: BER vs Eb/N0 – TM Modulations

## Slide 15: Technique
Adjust source EIRP, modulation, polarization, and bit rate to achieve a known Eb/N0
  Example:
    Eb/N0 of 11.4 dB yields BER = 10-5 (DQM = 5) in SOQPSK
  Solve for G/T
This is the G/T while the antenna is looking at the QSight™
NOT the same as G/T using “coldest sky”

## Slide 16: Does it Match Solar Method?
No. And it shouldn’t!
Solar method built to get the “best” G/T
  Find the “coldest” sky  Get the best G/T
  Derived G/T will vary with sky temperature
  Assumes solar flux is accurately interpolated
Signal source method yields the G/T while looking at the source
  Result may still vary with sky temperature
  This effect is “real”, belongs in the link budget

## Slide 17: Numerical Results

## Slide 18: Solves Several Problems
Eliminates solar flux estimation errors
Eliminates “cold sky optimism”
Works for small antennas
  No need to guess at tiny Y factors
Works for multi-beam antennas
  One source for each beam, spatially separated
  Get simultaneous G/T on all beams

## Slide 19: Proper Execution Requires…
Calibration of EIRP
  Currently accurate from -40 dBm to +40 dBm
  Extending the low end would be helpful
  Need to get down to target BER (or DQM)
Calibration of BER vs. Eb/N0
  Accurate DQM is helpful
Need sufficient distance
  Fraunhofer distance is a minimum

## Slide 20: Conclusions
G/T of a complete ground station can be found without using the sun
Yields meaningful G/T values
  Value belongs in your link budget
Works for any antenna
  Regardless of size
Works for multi-beam antennas
  Yields simultaneous G/T on all beams

## Slide 21: Backup Slides
In case of questions

## Slide 22: Start with a Link Budget
  Pt = Transmitter power
  Lt = Transmit cable losses
  Gt = Transmit antenna gain
  Gr = Receive antenna gain
  = Carrier wavelength
  D = Path length
  k = Boltzmann’s constant
  T = Receiving system temperature
  Rb = Information bit rate
determines BER performance…

## Slide 23: Typical Solar Flux Data

## Slide 24: time of day

## Slide 25: time of year - Monthly averages of solar 10.7 cm flux
Chart: Monthly Flux - 2004 to 2022

## Slide 26: years

## Slide 27: interpolation algorithms

## Slide 28: interpolation algorithms

## Slide 29: Error compared with linear fit CelesCoord.exe Program G/T calculator
Chart: Solar Flux Interpolation (3rd Order)
| 1425 | 2695 | 4995 | 8800
 | 60 | 74 | 120 | 237
Celest X-band SF linear fit
Celest S-band SF linear fit
Results in +0.2 dB error in G/T at 4401.5 MHz